This tool is organised by department and provides a range of typical procedures that may be encountered in that department. Probe use and reprocessing requirements based on current editions of local, regional and global guidelines are presented as a decision-making algorithm.1-7 The flow chart can be printed out and displayed throughout the office and procedure rooms and used as a quick reference chart for healthcare workers to determine whether a practice is compliant with available guidelines.

This tool can also be used in conjunction with Tool 1: Part B – Profile to determine whether policy and practice comply with best practice recommendations.

Table of Contents

Guidelines	2

Algorithm template	4

Anesthetics	5

Emergency	6

Cardiology	7

Gastroenterology	8

Obgyn/MFM	9

Musculoskeletal	10

Nephrology	11

Neurology	12

Oncology	13

Pulmonology	14

Urology	16

Vascular	17

References	18

# Guidelines

The following tool is based on the Spaulding Classification system that is referenced in U.K., European and global ultrasound guidelines.1-6 Local regulatory authorities could refer to these documents for accreditation of health care facilities.

The Spaulding Classification system dictates the level of disinfection or sterilisation a medical device should be subjected to based on the degree of infection risk involved with its use and contact sites with patients.11 The following definitions below are drawn from local and regional guidelines:1-7

Non-critical devices – only contacts intact patient skin, not mucous membranes. If decontamination is necessary, disinfect with compatible low or intermediate level disinfectant after cleaning.

Semi-critical devices – contacts intact mucous membranes or non-intact skin, such as skin broken from the insertion of a vascular device or complex wounds.1-6 Probes include a transoesophageal echocardiogram, transrectal ultrasound and transvaginal probes.2-6

A high-level chemical or thermal disinfection process that is bactericidal, mycobactericidal, fungicidal and virucidal with a small number of spores permitted must be used after cleaning.2-4 If this process includes equipment, it must be validated and also maintained following relevant standards and manufacturer’s instructions.2-4 HLD is recommended regardless of sheath use.1-6

Critical devices – probes contacting sterile tissues or blood, includes ultrasound probes used in sterile body cavities, intraoperative, biopsy and puncture techniques. The critical devices are not defined in the HSE Ireland and NHS Scotland guidelines though it is referenced in SCoR/BMUS, ECMUS and WFUMB guidance.

Critical devices require sterilisation. However, some sterilising methods can cause damage to ultrasound probes, and the SCoR/BMUS and ESR guidelines propose an alternative HLD combined with the use of a sterile sheath.1,4

Note: The ESR suggested in 2017 simplifying the classification by moving the semi-critical devices into the critical category. This toolkit does not implement this proposal and keeps three distinct categories in alignment with the U.K./Ireland recommendations. However, while performing any risk analysis, it is advisable to consider the fragility of mucous membranes and the risk of micro-trauma.

The following procedure charts are based on these recommendations and are divided by department. Typical procedures have been provided as examples based on interpretation of the Spaulding Classification. Some procedures may fall into other Spaulding Classification categories depending on the specific clinical situation and which tissues are contacted. Print these charts and place them in treatment rooms to guide best practice ultrasound reprocessing and use for patient safety. Any deviation from best practice recommendations needs to be supported with clinical evidence and thorough risk profiling with a multidisciplinary assessment team. This process needs to be documented and reviewed according to hospital policy and procedure.

Double click to type here

Type here

Type here

Type here

Algorithm template

Anesthetics

# Emergency

# Cardiology

Gastroenterology

Obgyn/MFM

Musculoskeletal

Nephrology

Neurology

Oncology

Pulmonology 

Radiology and Interventional Radiology

# Urology

# Vascular

# References

- Society and College of Radiographers and British Medical Ultrasound Society (2017). "Guidelines For Professional Ultrasound Practice." 127.

- Health Service Executive (HSE) Quality Improvement Division (2017). HSE Guidance for Decontamination of Semi‐critical Ultrasound Probes; Semi‐invasive and Non‐invasive Ultrasound Probes. Document: QPSD-GL-028-1.

- Health Protection Scotland (HPS) (2016). NHS Scotland Guidance for Decontamination of Semi-Critical Ultrasound Probes; Semi-invasive and Non-invasive Ultrasound Probes.

- Nyhsen, C. M., et al. (2017). "Infection prevention and control in ultrasound - best practice recommendations from the European Society of Radiology Ultrasound Working Group." Insights Imaging 8(6): 523-535.

- Kollmann, C. and K. Salvesen (2017). "Best Practice recommendations for cleaning and disinfection of ultrasound transducers whilst maintaining transducer integrity." European Committee for Medical Ultrasound Safety (ECMUS).

- Abramowicz, J. S., et al. (2017). "Guidelines for Cleaning Transvaginal Ultrasound Transducers Between Patients." Ultrasound Med Biol 43(5): 1076-1079.

- Department of Health (U.K.) (2016). Health Technical Memorandum 01-06 Part C Operational management. Department of Health (U.K.), Department of Health (U.K.),. C: 72.

- Department of Health (U.K.) (2016). Health Technical Memorandum 01-06 Part D Validation and verification (including storage/drying cabinets). Department of Health (U.K.): 49.

- Welsh Health Technical Memorandum (WHTM) (2014). WHTM 01-06 - Decontamination of flexible endoscopes Part C: Operational management, NHS Wales Shared Services Partnership – Specialist Estates Services: 74.

- Rutala, W. A., et al. (2008). Guideline for Disinfection and Sterilisation in Healthcare Facilities (Updated 15 February 2017).

- Spaulding EH (1968). Chemical disinfection of medical and surgical materials. Disinfection, sterilisation, and preservation. Lawrence C, Block SS. Philadelphia (P.A.), Lea & Febiger: 517-531.

- Masood J, Voulgaris S, Awogu O, Younis C, Ball AJ, Carr TW. Condom perforation during transrectal ultrasound-guided (TRUS) prostate biopsies: a potential infection risk. International urology and nephrology. 2007;39(4):1121-4.

- Amis S, Ruddy M, Kibbler CC, Economides DL, MacLean AB. Assessment of condoms as probe covers for transvaginal sonography. J Clin Ultrasound. 2000;28(6):295-8.

- Milki AA, Fisch JD. Vaginal ultrasound probe cover leakage: implications for patient care. Fertil Steril. 1998;69(3):409-11.

- Storment JM, Monga M, Blanco JD. Ineffectiveness of latex condoms in preventing contamination of the transvaginal ultrasound transducer head. South Med J. 1997;90(2):206-8.

- Rooks VJ, Yancey MK, Elg SA, Brueske L. Comparison of probe sheaths for endovaginal sonography. Obstet Gynecol. 1996;87(1):27-9.

- Hignett M, Claman P. High rates of perforation are found in endovaginal ultrasound probe covers before and after oocyte retrieval for in vitro fertilisation-embryo transfer. J Assist Reprod Genet. 1995;12(9):606-9.

- Fritz S, Hust MH, Ochs C, Gratwohl I, Staiger M, Braun B. Use of a latex cover sheath for transesophageal echocardiography (TEE) instead of regular disinfection of the echoscope? Clinical Cardiology. 1993;16(10):737-40.

- Abdelfattah R, Aljumaah S, Alqahtani A, Althawadi S, Barron I, Almofada S. Outbreak of Burkholderia cepacia bacteraemia in a tertiary care centre due to contaminated ultrasound probe gel. Journal of Hospital Infection. 2018;98(3):289-94.

- Shaban RZ, Maloney S, Gerrard J, Collignon P, Macbeth D, Cruickshank M, et al. Outbreak of healthcare-associated Burkholderia cenocepacia bacteremia and infection attributed to contaminated sterile gel used for central line insertion under ultrasound guidance and other procedures. American journal of infection control. 2017;45(9):954-58.

- Cheng A, Sheng WH, Huang YC, Sun HY, Tsai YT, Chen ML, et al. Prolonged postprocedural outbreak of Mycobacterium massiliense infections associated with ultrasound transmission gel. Clinical Microbiology and Infection. 2016;22(4):382.e1-.e11.

- Ofstead CL, Wetzler HP, Snyder AK, Horton RA. Endoscope reprocessing methods: a prospective study on the impact of human factors and automation. Gastroenterology nursing: the official journal of the Society of Gastroenterology Nurses and Associates. 2010;33(4):304-11.

- HSE Ireland - Guidance on the Implementation of the Health and Safety Competent Authority—Exemption 03/2016 on the Carriage of Uncleaned Reusable Invasive Medical Devices by Road (QPSD‐GL‐021‐1)

- Public Health England. Good infection prevention practice: using ultrasound gel. 2021; GW-1865

- Health Technical Memorandum (HTM) 01-01 on the management and decontamination of surgical instruments (medical devices) used in acute care. Part A: Management and provision – July 2016

- HSE Standards and Recommended Practices for Central Decontamination Units QPSD-D-003-2.1 V2.1. Part 3 Recommended Practices, 13.4 Procedure Pg 92.

- MHRA Medical Device Alert - Reusable transoesophageal echocardiography, transvaginal and transrectal ultrasound probes (transducers).MDA/2012/037 28 June 2012

- REGULATION (E.U.) 2017/745 OF THE EUROPEAN PARLIAMENT AND OF THE COUNCIL

of 5 April 2017 on medical devices, amending Directive 2001/83/E.C., Regulation (E.C.) No 178/2002 and Regulation (E.C.) No 1223/2009 and repealing Council Directives 90/385/EEC and 93/42/EEC.

- The Carriage of Dangerous Goods (Amendment) Regulations 2019.